A new approach to determining the hip rotation profile

from clinical gait analysis data

Richard Baker*, Laura Finney†,  John Orr†

* Musgrave Park Hospital, Belfast

† School of Mechanical and Manufacturing Engineering, The Queen’s University of Belfast

Corresponding author:	Richard Baker

Gait Analysis Service

Musgrave Park Hospital

Stockman’s Lane

Belfast		BT9 7JB

Northern Ireland

UK

Tel:	+44 1232 669501 ext 2155

Fax:	+44 1232 683816

e-mail:	richard.baker@greenpark.n-i.nhs.uk

Classification code:		2330

Keywords:	Three-dimensional gait analysis, Biomechanical Modelling, Hip rotation profiles

A new approach to determining the hip rotation profile

from clinical gait analysis data

Richard Baker*, Laura Finney†, John Orr†

* Musgrave Park Hospital, Belfast

† School of Mechanical and Manufacturing Engineering, The Queen’s University of Belfast

Abstract

Conventional models for determining joint rotation angles from marker positions as part of three-dimensional clinical gait analysis  are susceptible to errors arising from mis-placement of the thigh markers. An analysis of idealised data reveals how the measured variables are affected by different angular offsets of the thigh marker from its true position. An artefact on the varus-valgus signal arising from the projection of true knee flexion onto a mal-aligned thigh segment axis is the most characteristic feature of this problem. If this is observed then the hip rotation profiles are also erroneous.

A technique is proposed to determine a correction factor which can be applied to gait data to correct for this mal-alignment. Its use is demonstrated on a single case study and a subjective assessment of its use on a cohort of 40 patients is reported. A detailed discussion of the assumptions on which the method is founded is included as well as guidelines as to when the technique is likely to be successful.

It is recommended that the technique is used as an aid to training staff in marker placement rather than as a routine retrospective correction of data for erroneously placed markers.

- Introduction

Modern gait analysis requires a biomechanical model to infer the positions of body segments from the measured positions of markers placed on the subject’s skin. Kadaba et al. (1990) and Davis et al. (1991) proposed similar models which are the most widely used today. These model the lower limbs as seven rigid segments (one pelvis, two each of thigh, shank, and foot) which are defined by thirteen markers. The segments are assumed to be linked by joints with three rotational degrees of freedom. The respective orientation of segments is described by Cardan angles as suggested by Grood et al. (1983). The rotation of the distal with respect to proximal segment is described by angles representing first flexion/extension then ab/adduction then internal/external rotation.

This model is represented diagrammatically in Figure 1. All axis systems are approximately orientated with x anterior, y lateral and z superior on the left side. The axis system of the pelvis is defined by the three markers over the left and right anterior superior iliac spines (LASI and RASI) and the point midway between the posterior superior iliac spines (SACR). The position of the hip joint (HJC) in the pelvic axis system is calculated on the basis of co-ordinates scaled to leg length and the distance from one anterior superior iliac spine (ASIS) to the other.

The y-z plane of the thigh axis system is defined to lie in the plane containing the hip joint centre, a marker on a wand placed on the thigh (THI) and one placed over the lateral epicondyle of the femur (KNE). The knee joint centre (KJC) is taken to be half the previously measured knee width from the femoral epicondyle marker in this plane. The shank axis system is defined similarly to the thigh but using the knee joint centre, tibial wand (TIB) and lateral malleolus markers (MAL).

Measurements of internal and external rotation at the hip joint are therefore determined by the positioning of the thigh wand. The difficulty of doing this accurately and repeatably in subjects exhibiting pathological gait patterns is well understood clinically (CAMARC II: Clinical Movement Analysis Protocol for Cerebral Palsy,1995) but has been little discussed in the academic literature. Cappozzo et al. (1995) have proposed the use of redundant information from an array of markers on the thigh to determine an averaged segment orientation. This paper outlines an alternative approach of incorporating knowledge of joint anatomy into the biomechanical model to determine hip internal and external rotation.

In clinical decision making, for children with cerebral palsy in particular, accurate data for hip rotation profile are important as many subjects will have femoral anteversion leading to internal hip rotation. One of the important decisions requiring reliable information is whether the child would benefit from derotation osteotomies of the proximal femur.

# 2	Analysis

In their original paper Kadaba et al. (1990) drew attention to the fact that misplacement of the thigh wand will lead to the orientation of the axis system in the thigh being incorrect. If the axis system is not correctly defined then  true flexion/extension of the knee will be under-estimated and a varus-valgus artefact (cross-talk) will be introduced. They did not, however, draw attention to the fact that misplacement of the thigh marker will also lead to errors in the measured position of the knee joint centre. If the marker is too far anterior then the measured knee joint centre will be posterior to the true position and vice versa. This further affects calculation of flexion/extension and varus/valgus at the knee and also that of hip flexion and ankle dorsiflexion.

In order to understand the effects of thigh marker misplacement a computer model (using BodyBuilder software from Oxford Metrics, UK) was generated in order to calculate the effect of introducing perturbations on ideal data. A pelvis, left thigh and left shank segment were defined, linked by hip and knee joints. With hip and knee in neutral position the pelvis, thigh and shank axis systems were all aligned. The shank could then be rotated to a variable degree about the flexion axis of the knee (parallel to the y-axis of the thigh and shank). The dimensions of the segments were taken from an adult male subject and  the positions of markers SACR, RASI, LASI, THI, KNE, TIB and ANK as depicted in Figure 1 were calculated as if they were all ideally placed. The position of the THI marker was then varied to represent misalignment of the thigh wands. Its position is described as a rotation of its position about an axis joining the hip joint centre and the KNE marker (thigh marker offset).  The output of knee flexion/extension, varus/valgus and internal rotation were then calculated. Figure 2 shows graphs of how these vary with thigh marker offset for knee flexion of –20 (hyperextension), 0, 20, 40 and 60 degrees.

The measured knee flexion (see Figure 2a) shows the combined effect of the change in position of the knee joint centre, which will exaggerate knee flexion if the thigh marker is shifted posteriorly (-ve) and underestimate it if the marker is shifted anteriorly (+ve), and the rotation of the thigh axis system, which gives a reduced reading of knee flexion dependent on the size of the thigh offset but not its direction.  The offset giving rise to maximum measured knee flexion depends on knee flexion angle but is of the order of      –20º to -40. The curve for 20 hyperextension is the rotation of 180º about the graph’s origin of the curve for 20 flexion as required by the symmetry of the model.

For small values of thigh offset the measured knee valgus (see Figure 2b) is affected very little by the changing position of the knee joint centre. This is because for small offsets the change in joint position is restricted to the sagittal plane. It is only for larger offsets that the centre moves much in the coronal plane and it is this which leads to the curvature of the graphs. Within the range in which this effect can be ignored the measured knee varus/valgus varies linearly with the thigh marker offset. This shows that the major factor operating here is the projection of the true knee flexion angle (ft) onto the offset thigh segment axes. This can be described by the equation

vmsin(ft)							(1)

The measured value of vm for all points on the curves for knee flexions of –20 ,20 ,40 and 60 in the range of thigh marker offset between –10 and +10, differs from that predicted by Equation (1) by less than 0.2. The graph for 20 hyperextension is the reflection of that for 20 flexion in the graph’s vertical axis, a consequence of geometrical symmetry.

Figure 2c shows that for the lower degrees of knee flexion/extension there is a linear relationship between thigh offset and measured knee rotation. The gradient is approximately –0.7. There is an almost exact 1:1 relationship between thigh marker offset and the  rotation of the thigh axis system but the shifting of the knee joint centre also affects the alignment of the tibial axis system and so the effect is less pronounced on the measured knee rotation. The relationship shows some non-linearity for larger degrees of knee flexion.

# 3	Method

Although limited external rotation of the knee joint is predicted to occur at full extension on the grounds of joint anatomy during able-bodied gait, this is rarely observed in clinical gait data. This is partly because full knee extension is rarely obtained during gait and partly because the effect is masked by measurement errors (e.g. skin movement artifacts). Thus, for subjects exhibiting medio-lateral stability of the knee on passive examination, the knee joint can be considered, to a good approximation, to rotate in a single plane (Chao et al. 1983). Under these conditions any variation in the measured varus-valgus angle is likely to be a consequence of misplacement of the thigh marker.

The degree to which a thigh marker is misaligned can be represented by the thigh marker offset. A kinematic model based on the marker positions and model described in the Introduction above was written using the BodyBuilder software package (Oxford Metrics, Oxford, UK). This differs from the model described in the Analysis above in that it works on real marker data captured during gait rather than on ideal data. The varus-valgus angulation throughout the gait cycle can then be calculated for the thigh marker in its real plane and in a plane offset from this by any degree of thigh marker offset. The angle which results in the varus-valgus curve illustrating minimum cross-talk from knee flexion can then be taken as the offset required to correctly realign the marker (thigh correction factor). In principle this could be used to adjust the real position of the marker but in practice the correction is applied to the data by the software. Consideration needs to be given, however, as to how to analyse the curves to determine minimum cross-talk.

For small angles of true knee valgus we can assume that the measured valgus(vm) is the sum of the true valgus (vt) and a function (vf) of true knee flexion (ft) and thigh rotation offset (

vm  =  vt  + vf(ft, )						(2)

and the RMS variation of the signal about its mean value (statistical variance) is given by:

var(vm) = var(vt) + 2.cov(vt, vf) + var(vf)

where cov(vt, vf) is the covariance of vt and vf.

Now, from Equation 1 above

vf  .sin(ft)  for –10° < < 10°

and therefore for a fixed thigh rotation offset

var(vm) var(vt) + 2.cov(vt, sin(ft)). + var(sin(ft)).2

This is a quadratic equation in  with a minimum value (determined by setting the differential with respect to  equal to zero) at:

θmin=−covvt,sinftvarsinft					(3)

When considering the experimental data the only difference is that the thigh rotation offset (’) is measured from the real plane of marker placement, and therefore differs from the angle for thigh rotation offset with respect to the ideal plane () by a fixed thigh correction factor (). If min is small (say less than 1) then the value of ’ at the minimum point on the curve (’min) is a good approximation to .

At this point vt  vm and ft  fm and thus estimates of cov(vt,sin)ft) and var(sin(ft)) can be calculated to test the assumption that min is small. Once the correction factor has been determined then the data are reprocessed with the correction applied and the usual graphs generated.

## 4	Results

Figure 3 shows how knee flexion and varus-valgus data vary with thigh marker offset for sample clinical gait analysis data. As predicted in the analysis section both curves are affected by the offset. The curves representing varus-valgus show clear cross-talk with the maximum variation in signal with offset occurring during swing when there is maximum knee flexion but with some also when there is maximum stance phase knee flexion.

Figure 4 shows the variation of the variance of the measured knee varus with thigh marker offset for the left and right leg of a diplegic child. The curve is indeed seen to be symmetrical about a minimum value and the shape of the curve approximates to the square relationship as predicted above  (fit of quadratic regression curve gives R2=0.9995).

Figure 5a and 5b represent the data for hip internal/external rotation before and after correction for thigh offset using this method. The patient was a child with cerebral palsy. Physical examination and evaluation of video tape suggested a pattern of symmetric diplegia. Before applying the thigh offset correction it can be seen that the gait analysis data record an asymmetric pattern of hip internal/external rotation with the maximum internal rotation recorded for the two sides differing by about 15°. This is attributed to inaccuracies in placement of the thigh wands. After correction for these a more symmetrical pattern is observed.

The technique has now been used on 40 children with cerebral palsy with various levels of walking ability. A subjective assessment by clinical staff suggested that the corrected data were more believable than the uncorrected data for 60% of limbs.

## 5	Discussion

In the absence of a “gold standard” for hip internal/external rotation profile during gait it is difficult to substantiate the claim that this technique gives more accurate results than other methods. It is the subjective impression of the clinical team working in our gait analysis laboratory, however, that the corrected data consistently give a closer relation to visual (and video) observation of hip rotation profiles than the uncorrected data.

Figure 2 clearly show that cross-talk occurs between knee flexion and varus-valgus when ideal data are perturbed and Figure 3 that the cross-talk can be present in real data. The close correspondence of the variation in the variance of the varus-valgus data with thigh marker offset between theory and practice illustrated in Figure 4 reinforces the validity of the overall method.

The method rests on four assumptions, all of which can be tested as part of the analytical process and this gives internal checks to the validity of the method.

The first assumption is that the thigh marker offset remains constant throughout a gait cycle. This will not be the case if there is  movement of the thigh wands. Thigh wands move as a consequence of an oscillation due to mechanical instability of the mounting following rapid accelerations of the thigh and also of the movement of the underlying muscle bulk in more muscular subjects. The oscillation is of a higher frequency than the underlying gait pattern and can be filtered out if required (our experience suggests that this is not necessary). Movement of the muscle bulk has the potential to give rise to an artefact of similar frequencies as the underlying signal. Both will have the effect of increasing the variance of the “true” knee varus-valgus signal. This can be calculated once the theoretical offset has been calculated and the method should be rejected if this value is too high.

The second assumption is that the true varus-valgus angulation is small and that the two terms in Equation 2 can be combined by simple addition. This is a reasonable assumption if the true valgus variation is of the order of a few degrees. The validity of this  will be reflected by the value of the variance of the “true” knee varus-valgus signal. Again this suggests that the method should be rejected if this value is too high.

The third assumption is that min is small. This will occur if the numerator in Equation 3 is small or if the denominator is large. The process of choosing the thigh correction factor to  minimise the variance of the varus-valgus signal tends to lead to low values of the covariance term in Equation 2. The denominator is essentially a measure of how large a flexion signal there is to produce the cross-talk that the method is minimising. Calculation of min  allows the combined effect to be assessed and the method should be rejected if this value is too high.

The final assumption is that the artefact is purely a result of cross-talk from the knee flexion signal. If an additive error term (), which is independent of  is included in Equation 1 then Equation 2  becomes modified to:

θmin=−covvt,sinft+covε,sinftvarsinft			(4)

The only effect of other sources of error on the technique is therefore if they have an appreciable covariance with sin(ft). If this is the case then the estimate of the minimum point on the curve will be invalid. This again suggests that the method should be rejected if min is too high. The only possible situation in which errors would go undetected is if the two terms in the numerator of  Equation 4 were large and of opposite sign. As there is no reason why the true varus-valgus signal and other errors should show such dependence on each other this is not considered likely.

Use of the technique suggests that it should not be used if var(vt) > 9 deg2 (ie if the standard deviation of the residual varus-valgus throughout the gait cycle is > 3) or if |min| > 1.

### 6	Conclusion

A technique has been proposed to determine a thigh correction factor to compensate for mis-alignment of thigh markers. A theoretical justification for the method has been presented and various predictions of this model are borne out by analysis of real data. Use of this technique appears to give gait data which are more consistent with clinical observations than standard models. Low values of residual variance obtained suggest that this is a valid approach to determining hip internal/external rotation profile.

The technique remains a correction for mis-placed thigh markers. A better solution to the problem is to ensure more accurate placement in the first place. The technique  can serve as a check on marker placement and may well have a more important role in the training of  gait analysis staff in marker placement and in assuring the quality of gait analysis data than in routine correction of clinical data.

References

CAMARC II, Clinical Movement Analysis Protocol for Cerebral Palsy, (1995). Technical Report, Project A-2002 C.E.C. Programme AIM.

Cappozzo, A., Catani, F., Della Croce, U., & Leardini, A. (1995).  Position and orientation in space of bones during movement: anatomical frame definition and determination.  Clinical Biomechanics 10, 171-178.

Chao, E.Y., Laughman, E., Schneider, E., & Stauffer, R.N. (1983).  Normative data of knee joint motion and ground reaction forces in adult level walking.  Journal of Biomechanics, 16, 219-233.

Davis, R.B., Ounpuu S., Tyburski D., & Gage J.R. (1991). A gait analysis data collection and reduction technique. Human Movement Science, 10,575-587

Grood, E.S., & Suntay, W.J. (1983). A joint coordinate system for the description of clinical motions: Application to the knee. Transactions of the American Society of Mechanical Engineering: Journal of Biomechanical Engineering, 105,136-144

Kadaba, M.P., Ramakrishnan, H.K., & Wootten, M.E. (1990).  Measurement of Lower Extremity Kinematics During Level Walking.  Journal of Orthopaedic Research, 8, 383-392.

Figure 1

Biomechanical model of lower limb

Figure 2a, b and c.

##### Variation of measured knee flexion, valgus and rotation with

thigh rotation offset for five values of true knee flexion (idealised data)

Figure 3a,b and c

Sample clinical gait data: Variation of knee flexion/extension and valgus/varus and internal/external hip rotation with thigh marker offset. Data range for able-bodied subjects displayed in grey.

###### Figure 4

Variation of variance of measured knee varus with thigh correction factor. Sample data for left and right side. Bold line shows least squares second order regression. R2 = 0.9995 suggests that a quadratic fit to the curve for the right hand side is appropriate.

Figure 5a and b

Hip rotation profiles for child with symmetric diplegia

Before and after applying thigh correction factor.